# Audit template to assess compliance of low radiation dose CT used in the Targeted Lung Health Check Scans

Descriptor:

This audit is intended to assess the quality of Computed Tomography (CT) scans of the thorax in the setting of Targeted Lung Health Check scans.

Background:

The NHSE Targeted Lung Health Check (TLHC) protocol recommends the implementation of standardised acquisition and reconstruction parameters for TLHC low dose CT scans, as described in the NHSE Standard Protocol[1] and Quality Assurance Standards[2].

## The Cycle

The standard:

All teams delivering NHSE TLHC programmes should use the standardised acquisition and reconstruction protocols for low dose CT as described in the protocol [1,2]. Involvement of the medical physics department is recommended to audit doses and investigate if median dose length product (DLP) consistently exceeds local DRL. In the absence of a national diagnostic reference level (NDRL) a maximum recommended local DRL is 74 mGy.cm, but it may be appropriate to set a lower value based on local setup. For average patients this will approximately equate to 2mSv.

Target:

Image reconstruction is standardised and used for any subsequent follow-up examinations where possible with emphasis on ensuring that slice thickness, reconstruction increment, reconstruction algorithm is identical

- Median DLP for a sample of patients does not exceed the local DRL.

- 100% of reconstructed slice thickness of ≤ 1.25mm

- 100% of reconstructed slice increment of ≤ 0.7mm

- 100% scans reconstructed with correct standard lung algorithm.

- 100% of scans with correct anatomical coverage (i.e. entire lung parenchyma).

## Assess local practice

Indicators:

- Median DLP (assess if this exceeds local DRL)

- Percentage of scans reconstructed with slice thickness ≤ 1.25mm

- Percentage of scans with  reconstructed slice increment of ≤ 0.7mm

- Percentage of scans correct reconstruction algorithm.

- Percentage of scans with correct anatomical coverage.

Data items to be collected:

Number of scans included:

a) Total number of scans

Scan acquisition/reconstruction parameters:

b) DLP

c) Slice thickness

d) Reconstruction increment

e) Reconstruction algorithm

f)  Area of interest covered

Suggested number:

This will vary according to the size of the department. 100 examinations is recommended

Suggestions for change if target not met:

- Present and discuss findings CT optimisation team meetings.

- Correct acquisition and reconstruction settings, if required.

- Discuss possible solutions for quality improvement and plan for a quality improvement cycle.

Resources:

- List of eligible scans [may require support from PACS administration staff]

- Data collection from PACS (and RIS if required)

- Analysis and interpretation of data [which can be radiographer led]

References:

- [1] Standard Protocol prepared for the Targeted Lung Health Checks Programme, NHSE https://www.england.nhs.uk/publication/targeted-screening-for-lung-cancer/

- [2] Quality Assurance Standards prepared for the Targeted Lung Health Checks Programme, NHSE https://www.england.nhs.uk/publication/targeted-screening-for-lung-cancer/

Submitted by:

Peter Atiiga

Co-authors:

Craig Moore

Tim Wood

Lizzie Barclay

Andrew Stephens

Oliver Byass

Nadia Jawad

Published Date:

Wednesday 15 February 2023

Last Reviewed:

Tuesday 19 October 2021